|  | PRESS RELEASE |
| --- | --- |

The SHOUHUDUN respirator does not meet the requirements of FFP2

(Prague, 21 October 2022) The Czech Trade Inspection Authority has found that a respirator on the Czech market does not meet the requirements for FFP2 respirators and cannot effectively protect the user from solid particles or liquid aerosols.

The Czech Trade Inspection Authority has discovered a model of personal respiratory protective equipment - respirator - on the domestic market that does not meet the requirements of Regulation (EU) 2016/425 of the European Parliament and of the Council on personal protective equipment and repealing Council Directive 89/686/EEC. This is a product that was offered as "KN95 FFP2 Protective Mask" at the time of the inspection. It is a white filtering half mask with an embedded wire for shaping around the nose. The product was sold in 5 pieces in a transparent plastic package printed with the following information: KN95, Self-suction filter respirator, non-medical, GB2626-2006 KN95 Respirator mask, product information in a foreign language and the name and address of the manufacturer - Gaomi Pengshun Labor Protection Products Co. Ltd., Gaomi Economic Development Zone, Shandong Province.

Two paper self-adhesive labels with English description are stuck on the plastic packaging. The first label is marked with the date of manufacture on the packaging or in the product packaging (Production date / Date of manufacture), Expiration date 2 years from the date of manufacture (Expiration date / Shelf life). For more information please visit: http://www.flexio.cz/respirator. EN - KN95 respirator with filtration class FFP2 (5pcs) designed for disposable wear and respiratory protection, CE certification, Degree of filtration ≥95%, Material: meltblown and non-woven fabric, Made in China. Distributor CZ: FLEXIO a.s., U Hellady 697/4, 140 00 Prague 4, 210904, CE marking, EAN 8595235913215. Instructions for use were provided on the packaging in the form of pictograms with text in English only. In the packaging of the respirators there is a paper leaflet with the title Certificate of Conformity in English language and Chinese characters, the date of manufacture is 16.04.2020, lifetime 2 years.

As the product was offered under the name "KN95 FFP2 Protective Mask", laboratory tests were carried out in an accredited laboratory to verify the total permeation, permeation through filter material and breathing resistance according to the requirements set out in Articles 7.9.1, 7.9.2 and 7.16 of EN 149:2001+A1:2009. The results of the tests carried out show that the product did not meet the requirements for total permeation and permeation through the filter material - sodium chloride test and paraffin oil test. The product therefore does not comply with the requirements of Regulation (EU) 2016/425 of the European Parliament and of the Council on personal protective equipment. Due to the fact that the filtering half mask has a higher overall permeation and penetration through the filter material, it does not meet the requirements set for FFP2 type respirators and cannot effectively protect the user from solid particles or liquid aerosols. The product also fails the requirements for marking, instructions and consumer information. Furthermore, the EU Declaration of Conformity confirming  has not been provided on the product having followed the conformity assessment procedure in accordance with the requirements of the above mentioned Regulation.

The Czech Trade Inspection Authority imposed a protective measure on the distributor of the product, which is a ban on the supply on the market and use for own use of personal protective equipment, respirator brand SHOUHUDUN, and the withdrawal of this product from the market and circulation. Information on the findings relating to the product was reported to the MIT contact point and published in the Safety Gate system.

Product information:

Product name:				Protective mask KN95 5 pack

Trademark:				SHOUHUDUN

Model type / number:			not provided

Production serie number: 		only manufacture date 16 April 2020 provided

Identification data on importer:		not provided

Identification data on manufacturer:	Gaomi Pengshun Labor Protection Products Co. Ltd., Gaomi

Economic Development Zone, Shandong Province